## Slide 1
Deployment Details
Testing Organization and Site Information | 
Testing organization (Name, Organization Type, Contact website / phone number / email) | U.S. Environmental Protection Agency - Office of Research and Development https://www.epa.gov/air-sensor-toolbox/evaluation-emerging-air-sensor-performance PI: Clements.Andrea@epa.gov 919-541-1364
Testing location (City, State; Latitude & Longitude) | (AIRS) Ambient Monitoring Innovative Research Station 111 TW Alexander Dr. RTP, NC 27713 35.889510N, -78.874572W
AQS site ID | 37 – 063 – 0099
Sampling timeframe | Group 1: 2019-08-01 12:11:00 - 2019-09-02 04:59:00
Time Series Plots
Performance Metrics
Scatter Plots: Comparison to FRM/FEM
Meteorological Conditions During Deployment
Meteorological Influence
Sensor Information |  |  |  | 
Manufacturer, model | Example Make Model |  |  | 
Device firmware version |  |  |  | 
Sampling time interval | 1 minute |  |  | 
Sensor serial numbers | Group 1 SN01 |  | Group 1 SN02 | Group 1 SN03
Issues encountered during deployment? |  | Group 1: No Issues |  | 
FRM/FEM Information | 
Manufacturer, model | T-API T640X at 16.67 LPM
Sampling time interval | 
Date of calibration | 
Date of flowrate verification check | 
Description, date(s) of maintenance activities | 
1
For evaluations with greater than three sensors, grouping individual sensor metrics into boxplots is recommended for displaying results. Note that this recommendation does not apply to metrics computed as a single value for all sensors over the whole evaluation group, such as RMSE, NRMSE, CV, and standard deviation.
Sensor-FRM/FEM Accuracy
Sensor-Sensor Precision
Number of 24-hr periods outside manufacture-listed temperature target criteria | Group 1: 0
Number of 24-hr periods outside manufacture-listed relative humidity target criteria | Group 1: 0
Range of FRM/FEM concentrations over duration of base test (μg/m3) | Group 1: 3.3-15.3 (1-hr), 4.9-11.0 (24-hr)
Number 24-hr periods in FRM/FEM monitor measurements with a goal concentration of 25 μg/m3 | Group 1: 0
Mean number of paired, normalized concentration and temperature values | Group 1: 757.0
M number of paired, normalized concentration and relative humidity values | Group 1: 757.0
Testing Report - PM25 Base Testing
Example Make Model
Deployment #1
U.S. Environmental Protection Agency
Office of Research and Development
PI: Clements.Andrea@epa.gov
919-541-1364
May 2021

## Slide 2
  Additional documentation may be attached or linked to digital versions alongside this report. Such documentation may include field reports and observations during the testing period, maintenance logs for sensors and FRM/FEM monitors, standard operating procedures, and other documentation relevant to this testing report (see below for examples).
Supplemental Information
Supplemental Documentation | Attached | Description & URL or file path to documentation
Field observations |  | 
Maintenance logs |  | 
Standard operating procedure(s) |  | 
Photos of equipment setup and testing |  | 
Product specifications sheet(s) |  | 
Product manual(s) |  | 
Deployment issues |  | 
Data storage and transmission method |  | 
Data correction approach |  | 
Data analysis/correction scripts and version |  | 
Air Monitoring Station QAPP |  | 
Summary of FRM/FEM monitor QC checks |  | 
Other documents |  | 
Testing Report - PM25 Base Testing
Example Make Model
Deployment #1
U.S. Environmental Protection Agency
Office of Research and Development
PI: Clements.Andrea@epa.gov
919-541-1364
May 2021

## Slide 3
Testing Report - PM25 Base Testing
Example Make Model
Deployment #1
U.S. Environmental Protection Agency
Office of Research and Development
PI: Clements.Andrea@epa.gov
919-541-1364
May 2021
 | Bias and Linearity |  |  |  |  |  | Data Quality |  |  | 
 | R^2 |  | Slope |  | Intercept (μg/m^3) |  | Uptime (%) |  | Number of paired sensor and FRM/FEM concentration pairs | 
 | 1-Hour 0 out of 3 meet target | 24-Hour 0 out of 3 meet target | 1-Hour 3 out of 3 meet target | 24-Hour 3 out of 3 meet target | 1-Hour 3 out of 3 meet target | 24-Hour 3 out of 3 meet target | 1-Hour 3 out of 3 meet target | 24-Hour 3 out of 3 meet target | 1-Hour | 24-Hour
Metric Target Range | ≥ 0.70 | ≥ 0.70 | 1.0 ± 0.35 | 1.0 ± 0.35 | -5 ≤ b ≤ 5 | -5 ≤ b ≤ 5 | 75%* | 75%* | - | -
Sensor SN01 | 0.50 | 0.60 | 0.79 | 0.86 | -1.52 | -2.11 | 97.37 | 93.75 | 739.00 | 30.00
Sensor SN02 | 0.53 | 0.63 | 0.80 | 0.92 | -1.59 | -2.56 | 97.64 | 93.75 | 741.00 | 30.00
Sensor SN03 | 0.53 | 0.54 | 0.72 | 0.84 | -1.57 | -2.46 | 97.64 | 93.75 | 741.00 | 30.00
Mean | 0.52 | 0.59 | 0.77 | 0.87 | -1.56 | -2.38 | 97.55 | 93.75 | 740.33 | 30.00
 | Error |  |  | 
 | RMSE (μg/m^3) |  | nRMSE (%) | 
 | 1-Hour Metric target met | 24-Hour Metric target met | 1-Hour Metric target not met | 24-Hour Metric target not met
Metric Target Range | ≤ 7 | ≤ 7 | ≤ 30 | ≤ 30
Deployment Value | 3.75 | 3.57 | 48.23 | 46.50
 | Precision (between collocated sensors) |  |  |  | Data Quality |  |  | 
 | CV (%) |  | Standard Deviation (μg/m^3) |  | Uptime (%) |  | Number of paired sensor and reference concentration pairs | 
 | 1-Hour Metric target met | 24-Hour Metric target met | 1-Hour Metric target met | 24-Hour Metric target met | 1-Hour Metric target met | 24-Hour Metric target met | 1-Hour | 24-Hour
Metric Target Range | ≤ 30 | ≤ 30 | ≤ 5 | ≤ 5 | 75%* | 75%* | - | -
Deployment Value | 13.09 | 7.09 | 0.58 | 0.31 | 97.55 | 93.75 | 736 | 30
Tabular Statistics - Group 1
Sensor-Reference Correlation
Sensor-Sensor Precision